# Preface

The Sutter Instrument Company has provided this National Instruments VI library for customers who wish to use Labview virtual instruments to run the Lambda DG-4 filter changer.  With the exception of “N.D. Adjust.VI”, these are essentially drivers, intended for use as subVIs by our customers who already have some experience running Labview and creating Labview VIs.  If you are new to Labview, please take the time to familiarize yourself with the Labview and then read these instructions in full, before implementing the VIs in this package.

Labview DG-4 Library					Version 1.0

The Labview DG-4 Library is a set of virtual instruments, or VIs, for controlling the Sutter Instruments Lambda DG-4 from National Instrument’s Labview graphical programming language.

Installing and Using DG4.LLB the First Time

- Connect your PC to the Dg-4 using the parallel port.  See Appendix D for instructions on wiring the DG-4 end of the parallel port cable.

2.	Labview 4.1 or 5.0 should be installed and working properly on a computer running Windows 95.

3.	Find the file “labview.ini”, which should be in the base installation directory for Labview, usually c:\labview.  This is an ASCII text file.  Edit the file using the DOS Edit command in a command prompt window, or other ASCII text editor of your choice.  An example of the contents of this file is:

[LabVIEW]

FPFont="0" 13

BDFont="0" 13

hierarchy.stateFlags=71

find.findTextFlags=31736

find.selectVIsFlags=28672

find.viListFlags=16

IconFont="0" 13

If the file does not have a line beginning “serialDevices”, add this line at the end of the file:

serialDevices="COM1;COM2;COM3;COM4;LPT1;LPT2"

This line sets Labview device numbers corresponding to the serial and parallel ports.  It does not matter whether you actually have 4 COM ports in your computer; the line above will work anyway.  The first port listed, COM1, is assigned to Labview port 0 and so on, so that LPT1 is assigned to port 4.  If you use this string in labview.ini then you should use Port Number 4 in the DG4 VIs such as Initialize DG4 and Send String to Parallel Printer.  Save the edited file and close the Edit program.

4.	If you are using Labview version 4.1, install the updated Labview serial port driver, SERPDRV, which is supplied on the DG4 distribution diskette.  The driver that ships with this version contains bugs that prevent proper serial and parallel port operation.  First save the existing c:\labview\SERPDRV file to a floppy disk.  Then copy the updated SERPDRV from the DG4 distribution diskette to c:\labview or wherever you installed Labview.  If you are using Labview 5.0 or newer do not replace the SERPDRV with one provided by Sutter Instrument Company.

5.	Copy DG4.LLB to a directory you have created on your hard disk, for example c:\labview\dg4.

6.	Connect your computer’s LPT1 parallel port to the supplied “DG4 Parallel Adapter” and then through a straight-through 25-pin cable to the parallel port on the DG4 rear panel.  A straight-through cable is one that connects pin 1 to pin 1, pin 2 to pin2, and so on for all 25 pins on the DB-25 connectors.

7.	Labview and do File, Open, c:\labview\dg4\dg4.llb\Initialize DG4.vi.

8.	On the Initialize DG4 front panel, confirm that “Port Number” is set to 4.

9.	Run the VI.  It should run and stop quickly, with no error.  Look at the DG4 display: it should now say “Input = Parallel”.  If this worked correctly, go to the next step.

10.	Open and run the VI “Test DG4 Communication.vi”.  The DG4 display should update its “X = Filter Number” field as the VI runs.

What if this doesn’t work?

If Initialize DG4 returned an error, it is probable the cause is failure to communicate over the parallel port cable.  See the Troubleshooting section in Appendix A below.

Using Labview

This document assumes you are familiar with Labview programming.  Several of the demo and diagnostic VIs will be helpful in learning how to use the DG4 library.  Examine the wiring diagrams of these VIs:

N.D. Adjust

Test DG4 Communication

Ring Buffer Demo

Labview Standard Error Handling

All VIs in the DG4 library use standard Labview input and output error clusters.  You can string several VIs together, connecting the error output of one VI to the error input of the next and connecting the last error output to an error handler.  This ensures sequential processing of each VI in the string and handles errors that may occur anywhere in the string.  See sequence 0 of N.D. Adjust.vi for an example of this programming technique.

Individual Function Descriptions

Initialize DG4 – This MUST be the first command sent to the DG4.  It opens the parallel port for communication, sets up internal data structures for use by other commands, and sends command 238, Transfer Control to Computer, to the DG4.

Display On – Turns the DG-4 display on or off.  Connect True or False to the “Display On” input.

This VI can also be used to update the display after other commands have executed – see sequence 2 of the N.D. Adjust.vi diagram.

Filter Change – Moves to any filter in the range 0 to 15.  Moves immediately if “Move on Strobe” input is false, else the DG-4 will wait for the next strobe pulse.

Load Ring Buffer – Loads a sequence of up to 64 filter numbers into ring buffer memory.  Note that only 32 values can be retained in non-volatile memory; if you load more than 32 the excess will be lost when you power down the DG-4.

Ring Buffer Trigger Mode – This command selects the type of trigger that will be used in running from ring buffer:  Strobe pulse, Video sync, or Video sync gated by strobe.  There is also an option to deselect all triggers.  One of the three trigger types must be selected in order to use Run From Ring Buffer.  Programming note: the input “Trigger Mode” is an enumerated control.  If you wish to call Ring Buffer Trigger Mode from a VI you create, you will need an enumerated constant that matches the control.  To create one, go to the Ring Buffer Trigger Mode diagram, right-click on the Trigger Mode terminal, and select “create constant”.  You may then drag the constant to your VI and set its value using the operating tool cursor.

Run From Ring Buffer – This command prepares the DG-4 to step through a sequence of filters that has been loaded into the ring buffer.  The DG-4 requires a Sync Pulse or a Strobe Pulse to actually change from one filter to the next.  See Ring Buffer Demo for an example of how to use the three ring buffer commands.

Run From Ring Buffer has one Boolean input: YES = Run from Ring Buffer, NO = End Run from Ring Buffer. NOTE: While End Run from Ring Buffer is included in the Labview driver package, the DG-4 currently does not respond to End Run. The DG-4 busy line only goes low for short periods of time and the Labview system cannot detect the very short “non-busy”.  The short low on the busy line is an intentional design on the part of Sutter to enable the ring buffer to run as fast as possible.

N. D. Adjust – Allows you to adjust the attenuation of a given filter number by controlling the output galvanometer value associated with that filter.  This is a high-level VI that is intended to run standalone; it is not callable from other VIs.  Enter the desired filter number in the Filter Number control, and then press the Run arrow.  The VI loops until you press the Stop&Save key.  As soon as you run the VI, the DG-4 display should change to read:

Filter No. X		Current Value = XXXX

As you press the numeric keys on the VI, +1, +10, -1, and –10, the Current Value will change.  When the value is satisfactory, press the Stop&Save button and the DG-4 will store the value.

Open Shutter – Opens or closes the shutter by sending True or False, respectively, to the Open Shutter control.

Turbo Blanking – Turns turbo-blanking mode on or off by sending True or False, respectively, to the Turbo Blanking On control.

Diagnostic VIs

The following VIs are for diagnostic or demonstration use.

Send String to Parallel Printer - This VI tests whether Labview can send data to the parallel port.  Connect a dot-matrix printer to LPT1 using a standard Centronics parallel cable (DO NOT use the DG-4 Parallel Adapter), set the correct Port Number, fill in some string you wish to send in the String to Send control, then run the VI.  If no errors occur, this confirms that Labview can write to the parallel port.  See the section Installing and Using DG4.LLB the First Time for instructions for setting up the Port Number.

Test DG4 Communication – This VI calls Initialize DG4 then loops continuously, calling Filter Change for filters 0 to 15.  Once you have verified parallel port operation using Send String to Parallel Printer, you should use this VI to verify DG-4 communication.  Note:  be sure you are using the DG-4 Parallel Adapter and a straight-through 25-conductor cable to connect the parallel port to the DG-4.

Ring Buffer Demo – A demonstration program that shows how to use the three ring buffer commands.

Low Level VIs

Send Command to DG4 – This and Initialize DG4 are the two lowest level VIs in the library.  They use Labview’s built-in commands Serial Port Init and Serial Port Write to send data to the serial or parallel port.  All other VIs in DG4.LLB funnel data through Send Command to DG4.  You will normally use the higher level VIs described above and not call Send Command to DG4 directly, but you may use it if you wish.  Remember that Initialize DG4 must be called first (once).  The DG-4 Command input must be an unsigned byte in the range 0 – 255.

Wait Short – Waits a short amount of time (60ms default) in order to give DG-4 commands time to execute.  Called by other VIs in the library.  The wait time is defined by the global value DG4_global_Short_Wait, in DG4_Globals.  This value is set by Initialize DG4.

DG4_Globals – Holds global parameter values that are used by other VIs in the library.  The values are set by Initialize DG4.

APPENDICES

Troubleshooting

Failure to communicate with the DG-4 is most likely caused by one of the reasons listed in the sections below.  There are also two diagnostic VIs that can help in troubleshooting.  The first is Send String to Parallel Printer – see its description above.  This VI will verify that Labview can send data out the parallel port, independent of the DG-4.  The second VI is Test DG4 Communication.  This should be your testbed for verifying DG-4 operation from Labview.

Incorrect Labview Port Number

As supplied, the DG-4 library assumes the parallel port is mapped to  Labview port number 4.

Find the file “labview.ini”, which should be in the base installation directory for Labview, usually c:\labview.  This is an ASCII text file.  Edit the file using the DOS Edit command in a command prompt window, or other ASCII text editor of your choice.  An example of the contents of this file is:

[LabVIEW]

FPFont="0" 13

BDFont="0" 13

hierarchy.stateFlags=71

find.findTextFlags=31736

find.selectVIsFlags=28672

find.viListFlags=16

find.nodeRsrcID=1050

saveFloaterLocations=True

editRecentPaths="D:\LABVIEW\SEGAL\Utility.llb;

IconFont="0" 13

If the file does not have a line beginning “serialDevices”, add this line at the end of the file:

serialDevices="COM1;COM2;COM3;COM4;LPT1;LPT2"

This line sets Labview device numbers corresponding to the serial and parallel ports.  It does not matter whether you actually have 4 COM ports in your computer; the line above will work anyway.  The first port listed, COM1, is assigned to Labview port 0 and so on, so that LPT1 is assigned to port 4.  If you use this string in labview.ini then you should use Port Number 4 in the DG-4 VIs such as Initialize DG4 and Send String to Parallel Printer.  Save the edited file and close the Edit program.  Close and restart Labview if it was running.

2. Parallel Port in Use by a Printer Driver

Even though no Windows program is actually printing to the printer port, it still may be "in use" as far as Windows is concerned.  Many recent printers, such as HP Deskjet and Laserjet printers, use complex printer drivers that check on printer status at all times.  These drivers tie up the port continuously, preventing access by Labview.  If you have one of these printers, try first deselecting the printer by making a simple dot-matrix printer the default.  For example, install an Epson MX-80 driver, make it the default printer, then shutdown and restart the computer.

If this doesn't work, remove the offending printer driver completely.  Go to Control Panel, Printers, right-click on the printer icon, and select "Delete".  Then shutdown and restart the computer.

If you do not wish to dedicate LPT1 to Labview's use, you may purchase a printer adapter card and install it in the computer.  Once properly installed it should appear as LPT2.  Make sure you use the correct Labview port number for LPT2.

3. Parallel Port Mode

The computer’s parallel port mode should be set to “compatible”, “standard”, or “unidirectional”.  This is set in the computer’s CMOS or BIOS Setup program, which is usually accessed by pressing the DEL or F1 key during bootup.  If the port is set to “extended” or “bi-directional” mode, the signals may not be compatible with the DG-4 requirements.

4. Incorrect Cable Connections

See Appendix D.

5. Incorrect SERPDRV Version

If you are using Labview 4.1, install the updated Labview serial port driver, SERPDRV, which is supplied on the DG-4 distribution diskette.  The driver that ships with Labview 4.1 contains bugs that prevent proper serial and parallel port operation.  First save the existing c:\labview\SERPDRV file to a floppy disk.  Then copy the updated SERPDRV from the DG-4 distribution diskette to c:\labview or wherever you installed Labview.  If you are using Labview 5.0 or newer do not replace this driver.

6. Failure to use Initialize DG4

The first command sent to the DG-4 MUST be from Initialize.vi.

Appendix B.  Error codes

1 or 37 - Failed to Open Port or Device Not Found.  "Initialize DG4.vi" failed to open the parallel port.  Possible causes: Port Number is incorrect for LPT port or LPT port is in use by another Windows program.

See Appendix A, sections 1 and 2.

2 - Filter number out of range.  Filter numbers must be in the range 0 . . 15.  When using the "Move on Strobe" function in Filter Change.vi, set the "Move on Strobe" control True but leave filter number in the range 0 . . 15.

3 - DG-4 command out of range.  All commands to the DG-4 must be unsigned bytes in the range 0 . . 255.

4 - Load Ring Buffer - either array too large or an individual filter number out of range.  Filter numbers must be in range 0-15 and the array cannot contain more than 64 elements.

37 - See #1, Failed to Open Port.

Appendix C.  Distribution diskette contents

DG4.LLB 	– Labview library file.

SERPDRV 	– Updated serial port driver for Labview 4.0 and 4.1.  Should have 36,020 bytes.

DG4.DOC 	– DG4 library documentation in Microsoft Word format (this document).

Appendix D.  Parallel cable wiring

The DG-4 was not initially designed to work with Labview.  Since Labview was designed to treat the parallel port as if it were connected to a printer, the parallel cable connecting an IBM PC parallel port to the DG-4 parallel input must be wired as follows:

Computer end					DG-4 end

DB-25 Male	Function	Direction	DB-25 Female		Comment

1		*STROBE	>		N/C

2		Bit 0		>		2

3		Bit 1		>		3

4		Bit 2		>		4

5		Bit 3		>		5

6		Bit 4		>		6

7		Bit 5		>		7

8		Bit 6		>		8

9		Bit 7		>		9

10		*ACK		<		16			*ACK tied high

11		BUSY		<		11

12		PE		<		25			PE grounded

13		SELECT	<		16			SELECT tied high

15		*ERROR	<		16			*ERROR tied high

16		*INIT		>		16

24		GROUND			24

25		GROUND			25

Note:  On the DG-4 end of the cable, signals *ACK, SELECT, and *ERROR are tied back to *INIT, which is normally in the high (+5V) state, thus pulling these three signals high.  Signal PE is tied to GROUND.  These signals must be tied to +5V and ground, respectively, in order for Labview to communicate with the DG-4.

There are two approaches to creating the cable.  The first is to wire a custom cable as shown above.

The second is to buy a standard straight-through 25-pin cable with a DB-25 Male on one end and a DB-25 Female on the other, then use an adapter box to obtain the required wiring for the signals that do not go straight through.  When using this approach, wire the adapter box as shown above and plug its male connector into the computer.  Attach the cable to the female side of the adapter box and run it to the DG-4.

A suitable male-to-female adapter box can be purchased from Digikey Corporation, 1-800-344-4539.  It is manufactured by Assman, stock number AE1039, $7.82.  A suitable 9-foot straight-through 25-pin cable is also available from Digikey, stock number AE1013, $8.58.