# Special Note to the Lambda 10-2 User

The Labview Virtual Instruments you have been provided with were originally intended for use with the Sutter Instrument Company’s Lambda DG-4.  As a result, you will find many of these VIs are non-essential.

Literature has been included explaining both the Labview DG-4 library software and the Lambda 10-2 controller.  This page is intended to answer obvious questions arising from the fact that the Labview manual does not refer to the Lambda 10-2 specifically, and the manual for the 10-2 makes no reference to it’s use with Labview software.

Although the individual manuals for each will give the user a much broader overview of the Lambda 10-2’s capabilities, and how the Labview software is intended to be used, the following is a brief explanation of the two Labview sub-VIs that you most likely be using first.

## Initialize DG-4

This MUST be the first command sent to the Lambda 10-2.  It opens the parallel port for communication, sets up internal data structures for use by other commands, and sends command 238, Transfer Control to Computer, to the Lambda 10-2.  As mentioned, this was designed to work with the DG-4, but works the same way with the Lambda 10-2.

## Send Command to DG-4

This will most likely be your most utilized VI.  Using this as a sub-VI, you can design the appropriate interface for the 10-2, making sure that the appropriate byte is sent to “V8” control labeled “DG-4 command” in this VI’s block diagram.

If you have established contact with the Lambda 10-2 using the Initialize DG-4 VI, you can open this VI and simply enter appropriate number in the box labeled “DG-4 command.”   These numbers correspond to 8-bit bytes that inform the controller what to do next.

For example, 0 moves wheel A to filter # 0, 1 moves wheel A to filter # 1, etc.  To change the filter in wheel B, 128 will move the wheel to filter # 0, 129 moves it to filter # 1, and so on.  Please refer to your Lambda 10-2 manual for a much more detailed explanation of how these specific numbers control your Lambda filter wheel.

As a sub-VI, this will be used in any interface you choose to design using Labview.